EL-USB Data Logger Instructions

Dear Customer,

Thank you for purchasing an EL-USB temperature data logger.

Before using the logger, you must first configure it using the EasyLog-USB software. This allows you to configure the logger and automatically download the measurement data to a Windows PC.

Download the free software via:
www.dataloggers.shop/Software

Getting Started

- First use: configure the logger via the EasyLog-USB software

- After each use: reprogram the logger using the same software before starting a new measurement cycle

- Note: battery life depends on the selected logging interval and the ambient temperature.
We recommend not setting the interval too short unless necessary. For most applications, one measurement every 15 minutes is sufficient.

- Ensure that the plastic cap is properly placed to prevent moisture from entering the logger.

Logger not starting – Battery passivation

If the logger does not start during first use, it may be due to battery passivation, a common phenomenon in lithium batteries.

Solutions:

- Firmly tap the battery a few times against a hard surface to break the passivation layer

- Leave the logger with the battery inserted in the USB port for at least 30 seconds

- Remove and reinsert the battery

These actions can help restore functionality.

Reading the Serial Number

You can find the logger’s serial number in the software under:
Options > Current Status

Battery Information (from manufacturer’s manual)

Replacement

- Replace the battery once a year or before logging critical data

- Only use 3.6V ½AA lithium-metal batteries

- The logger will not lose data when the battery is removed or discharged

- However, logging will stop and only resume after battery replacement and a restart using the EL-WIN-USB software

- Remove the logger from the PC before changing the battery

- Extended connection to a USB port may shorten battery life

New battery not working – Passivation

Unused lithium-metal batteries may develop a non-conductive internal layer, preventing discharge during storage (passivation).
When installed, this may cause a temporary voltage drop and reset the logger.

To resolve this:

- Insert the battery and leave the logger connected to the USB port for 30 seconds

- Remove and reinsert the battery

- Firmly tapping the battery on a hard surface often yields the best results

This does not affect the total battery life.

Warning

Handle lithium-metal batteries with care and follow all safety instructions.
Dispose of used batteries according to local regulations.

Happy logging!
Tempro
www.dataloggers.shop
✉ mail@tempro.be